## Slide 1
OISL — Experimental Test Setup
Hardware & Software Chain Design
Inter-Satellite Laser Communication · Optical Component Verification
Dina | KAIST – Aerospace Engineering | PhD Research
Optical Bench
RF & Elec. Instruments
Signal Chain
DAQ & FPGA
Software Stack
Notes: This deck covers the complete experimental test setup for verifying OISL optical components, from hardware bench equipment through signal chain to the software analysis stack. Matches the component-level design topics from the previous deck.

## Slide 2
TEST PHILOSOPHY
Experimental Strategy & Test Hierarchy
Component-level
Characterize each optical element individually (laser, FSM, detector, telescope). Isolate variables.
Subsystem-level
Validate TX chain, RX chain, and PAT loop independently before integration.
System-level
End-to-end link test with range-equivalent path loss. Full PAT acquisition and data transmission.
Environmental
Thermal vacuum, vibration, and radiation qualification per ECSS / NASA standards.
Overall Test Chain Flow
DUT
Setup
Signal
Stimulus
Optical
Path
Detection
& DAQ
Software
Analysis
Pass/Fail
Verdict
Hardware / DUT
RF & Elec. Instruments
Optical path elements
DAQ / FPGA / digitizer
Software stack
Notes: Testing follows a V-model hierarchy: component → subsystem → system → environmental. Each level has specific instruments, setups, and pass/fail criteria linked back to the link budget and WFE budget from the design deck.

## Slide 3
HARDWARE — OPTICAL BENCH
Optical Bench & Beam Path Equipment
Isolation table
Newport RS2000 / TMC optical table, 4′×8′, active pneumatic isolation. Cuts > 10 Hz floor vibration. Essential for μrad-level alignment.
Precision mounts
6-DOF kinematic mounts (Thorlabs Polaris, Newport 562), sub-μrad tip-tilt, 50 nm positional resolution. Holds telescope, FSM, detector.
Beam steering
2-axis piezo mirror mount (Physik Instrumente S-330 or Newport AG-M100N). Emulates satellite jitter. Drive: ±1 mrad, 1 kHz BW.
Beam expander / collimator
Galilean expander (Thorlabs BE10M or custom). Verified with Shack-Hartmann WFS. Expansion ratio 5–20× depending on DUT aperture.
Polarization optics
PBS cube + λ/2 + λ/4 waveplates (Thorlabs WPQ / WPH achromatic). Set and monitor polarization state. Critical for MZM input.
Neutral density filters
Calibrated ND filter set (OD 0.3–6.0) for attenuating transmit power to safe detector level. Also simulates link path loss.
Specialty Optical Elements
Fizeau interferometer
Zygo GPI or 4D PhaseCam. Measures surface figure of mirrors/lenses in-situ. λ/100 measurement accuracy. Used before + after vibration test.
Shack-Hartmann WFS
Thorlabs WFS300-14AR. 150-lenslet array, 1550 nm. Measures transmitted wavefront error of full beam path. Real-time Zernike modal decomposition.
Beam profiler
Spiricon M2-200 or DataRay WinCamD. Measures beam quality M², 1/e² radius, centroid position. Verify Gaussian profile at aperture exit.
Power meter + integrating sphere
Thorlabs PM100D + S145C InGaAs head. Calibrated ±3% absolute. Measures η_tx, η_rx, and insertion loss of each element.
Delay-line emulator
Fiber spool (SMF-28) of calibrated length to add time-of-flight delay equivalent to target range. Typically 1–10 km fiber = 5–50 μs delay.
Notes: The optical table is the single most important piece of infrastructure — without sub-Hz vibration isolation you cannot achieve μrad-level alignment stability. The Shack-Hartmann WFS is the real-time WFE monitor; the Fizeau gives higher-accuracy static measurements.

## Slide 4
HARDWARE — TX CHAIN INSTRUMENTS
Laser Source, Modulator & Amplifier Test Equipment
Tunable Laser Source (TLS)
Make/model:
Keysight 81600B or Santec TSL-710
Tuning range:
1480–1640 nm (covers 1550 nm ISL band)
Linewidth:
< 100 kHz — matches DFB seed spec
Output power:
Up to +13 dBm (20 mW) — seed level
Use:
Replaces DFB seed for characterization; known linewidth reference
Optical Spectrum Analyzer (OSA)
Make/model:
Yokogawa AQ6370D or ANDO AQ6317
Wavelength range:
600–1700 nm
Resolution:
0.02 nm — resolves EDFA ASE noise floor
Dynamic range:
> 70 dB — measures OSNR
Use:
Verify EDFA gain flatness, ASE noise, OSNR of modulated signal
Vector Network Analyzer (VNA)
Keysight E5063A or R&S ZVL13
S-parameter measurement of modulator electro-optic BW
Measure EO bandwidth of MZM (target > 40 GHz)
Also: FSM PZT electromechanical frequency response
Arbitrary Waveform Generator
Keysight M8195A or Tektronix AWG70000A
Generates PRBS data pattern (PRBS-31)
QPSK/DPSK IQ baseband drive for MZM
Variable symbol rate 1–40 Gbaud, calibrated amplitude
EDFA Characterization
EDFA test: OSA + calibrated attenuator + power meter
Gain vs. input power curve (saturation characterization)
Noise figure measurement (NF) via optical noise loading
Target: NF < 5 dB, gain > 20 dB over 1530–1565 nm
Notes: The TLS replaces the DFB during characterization to allow wavelength scanning and known linewidth testing. The OSA is essential for verifying OSNR — the EDFA adds amplified spontaneous emission (ASE) noise that must be within the link budget.

## Slide 5
HARDWARE — RX CHAIN INSTRUMENTS
Detector, Receiver & BER Test Equipment
APD Detector Characterization
Bias supply:
Keithley 2400 SMU — sweep reverse bias 0→−70V, measure dark current vs. APD multiplication M
Responsivity:
Calibrated power meter (Thorlabs S148C) + TLS sweep → A/W responsivity curve vs. wavelength
Bandwidth:
ESA (Keysight N9010B) + modulated laser source → measure −3 dB electrical bandwidth vs. bias voltage
NEP:
Low-noise TIA (FEMTO DHPCA-100) + lock-in amplifier (SR830) → noise-equivalent power floor measurement
Linearity:
Attenuator sweep + power meter → confirm linear range bounds (avoid saturation regime)
BER & Sensitivity Measurement
BERT:
Anritsu MP1800A or Ixia XGEM — pattern generator + error detector. Generates PRBS-31, counts bit errors in real time.
Attenuator:
JDS Uniphase / Agilent 81576A variable optical attenuator (VOA). Sweeps RX input power from −25 to −55 dBm.
Clock recovery:
Tektronix CKRX80G or integrated CDR in BERT. Recovers data clock from received waveform.
Oscilloscope:
Keysight DSA-X 86100D sampling scope — eye diagram, Q-factor, jitter decomposition. Bandwidth: 80 GHz.
BER vs. P_rx:
Sweep VOA: record BER at each power level. Plot waterfall curve. Read sensitivity at BER = 10⁻¹²
Electrical Spectrum Analyzer (ESA)
Keysight N9010B EXA, 10 Hz–26.5 GHz. Measures RIN of laser, phase noise of clock, and spurious tones in demodulated signal.
Optical Power Meter
Thorlabs PM100D + InGaAs head (S148C). Absolute calibration traceable to NIST. Logs time-series power during link test. ±3% accuracy.
Sampling Oscilloscope
Tektronix DPO77002SX, 70 GHz analog BW. Eye diagram, jitter analysis, rise/fall time. Connected post-TIA to verify waveform integrity.
Notes: BER waterfall curve is the key receiver figure of merit. The sensitivity at BER=10E-12 is compared against the link budget's receiver sensitivity requirement. Eye diagram closure (Q-factor) reveals if degradation is from noise or inter-symbol interference.

## Slide 6
HARDWARE — PAT TEST SETUP
Pointing, Acquisition & Tracking Test Hardware
Jitter & Disturbance Emulator
2-axis piezo mirror (Physik Instrumente S-330.2SL) driven by arbitrary waveform generator. Injects controlled angular jitter to emulate satellite reaction wheel disturbances. Amplitude: 0–2 mrad, Freq: 1–1000 Hz.
4-Quadrant Detector (4-QD)
Hamamatsu G6849-01 or OSI QP50-6SD InGaAs 4-QD. Bandwidth: > 100 kHz. Measures beacon spot centroid as differential photocurrent. Connected to analog summing/differencing circuit.
High-speed CMOS Camera
Xenics Bobcat-1.7-320 or Allied Vision Goldeye G-032. 320×256 InGaAs pixels, 400+ fps at full frame. Used for wide-field beacon acquisition phase. FOV: 5°×5° with 25 mm FL lens.
FSM + High-Voltage Driver
PI S-330 FSM + E-727 piezo controller. Analog input ±10V → ±5 mrad. Controller provides closed-loop capacitive sensor feedback. Position resolution < 0.1 μrad. BW: > 1 kHz.
Coarse Gimbal (2-axis)
Physik Instrumente U-651 motorized rotation stage or Newport RV240CC. ±180° range, 0.001° resolution. Driven by motion controller (Newport XPS or PI C-884). Emulates satellite coarse pointing.
Real-Time Controller (RTC)
National Instruments cRIO-9049 with FPGA + LabVIEW RT. Runs PAT control loop at 50–100 kHz update rate. Implements PID + notch filters + Kalman estimator in deterministic FPGA fabric.
Notes: The jitter emulator is critical — it injects realistic reaction wheel spectral content so the PAT loop can be tuned against real disturbance signatures. The cRIO runs the deterministic control loop; non-deterministic tasks (logging, UI) run on a separate host PC.

## Slide 7
HARDWARE — DAQ, FPGA & DIGITIZERS
Data Acquisition & Real-Time Processing Hardware
Core Real-Time Platform: NI cRIO / Xilinx FPGA
FPGA fabric:
Xilinx Kintex-7 325T (cRIO-9049) or Virtex-7 690T for heavy DSP
Loop rate:
Control loop: 100 kHz deterministic; DAQ streaming: up to 250 MSa/s
I/O:
±10V analog I/O (NI 9220 + 9264), 32-bit DIO, counter/timers for encoder readout
DRAM buffer:
4 GB DDR4 on RIO board for burst capture during acquisition transients
High-Speed Digitizer
NI PXIe-5186 or Keysight U5303A
12-bit, 3 GSa/s, 2 ch
Captures TIA output waveforms
FFT noise floor analysis
Synchronized to AWG trigger via PXI backplane
Lock-in Amplifier / PSD
Stanford Research SR860 DSP
1 MHz BW, 24-bit ADC
Measures signal at modulation freq
NEP floor characterization of APD
Phase-sensitive WFE measurement
Timing & Synchronization
Keysight 53230A universal counter
12-digit freq resolution
Time-interval measurements (PAT timing)
GPS-disciplined 10 MHz reference
Phase-locks AWG, OSC, cRIO to common timebase
Thermal Monitoring DAQ
NI 9213 thermocouple module (16 ch)
Type-K TCs on optics, EDFA, detector
Logs thermal drift vs. WFE correlation
Triggers thermal control loop
ECSS-compliant telemetry archive
Notes: The PXI backplane is the backbone — all instruments share a common 10 MHz reference and trigger bus so every capture is time-correlated. The FPGA closes the control loop deterministically; the host PC does logging and analysis asynchronously.

## Slide 8
SOFTWARE STACK — OVERVIEW
Software Layers: From FPGA to Analysis
Layer 5 — Analysis & Publication
Python / MATLAB / Julia · NumPy, SciPy, matplotlib, pandas
Layer 4 — Test Automation & Orchestration
LabVIEW TestStand / Python pytest · SCPI instrument drivers · CI pipeline
Layer 3 — Real-Time Control (Host PC)
LabVIEW RT / MATLAB Simulink RT · PAT state machine · Kalman filter SW model
Layer 2 — FPGA Deterministic Loop
LabVIEW FPGA / VHDL / Vivado · PID + notch filters · 100 kHz control cycle
Layer 1 — Hardware Abstraction & Drivers
NI-VISA / NI-DAQmx / IVI-C drivers · SCPI over GPIB/USB/LAN · custom FPGA IP cores
FPGA code (Layer 2) is compiled with Vivado synthesis — not interpreted. Control loop timing is deterministic to ±10 ns.
Python orchestration layer communicates via SCPI over LAN to all instruments simultaneously — parallel instrument sweeps reduce test time 3–5×.
Notes: The layered architecture is critical for maintainability. Layer 1-2 are hardware-specific and rarely change. Layer 3-4 contain the test logic. Layer 5 is purely post-processing. Different engineers can work on different layers without stepping on each other.

## Slide 9
SOFTWARE — FPGA & REAL-TIME CONTROL
FPGA Control Loop & Real-Time Software
FPGA Control Loop (100 kHz)
1. ADC read (10 μs)
Read 4-QD differential photocurrent (X,Y) from NI 9220 ADC at 100 kHz. Compute centroid error: ε_x = (A+B−C−D)/(A+B+C+D).
2. Error filter (1 μs)
Apply digital bandpass (100–800 Hz) to separate tracking signal from DC offset and high-frequency noise. IIR second-order Butterworth.
3. PID compute (2 μs)
K_p·ε + K_i·∫ε·dt + K_d·dε/dt. Implemented as velocity-form PID to avoid integrator windup on actuator saturation.
4. Notch filters (2 μs)
Series of 2nd-order IIR notch sections. Frequencies: plant resonance (ω_n), RW fundamental (f_RW), 2nd harmonic (2f_RW), 3rd (3f_RW).
5. Kalman predict (3 μs)
State: [θ, θ̇, d_dist]. Predict step: x̂_k|k−1 = F·x̂_k−1. Disturbance feed-forward: u_ff = −K_ff·d̂_dist.
6. DAC write (2 μs)
Write FSM drive voltage to PI E-727 controller via NI 9264 DAC. Saturation limit: ±9.5V (prevents actuator overtravel).
Kalman Filter — State Space Model
State vector x:
x = [θ, θ̇, d_dist]ᵀ — mirror angle, angular velocity, disturbance force estimate
State transition F:
Discretized at T_s = 10 μs. F = expm(A·Ts), A from plant model P(s) state-space realization.
Process noise Q:
Tuning knob: models unmodeled dynamics. High Q → tracks fast disturbances, more noise. Set via empirical Allan variance of jitter signal.
Measurement noise R:
R = σ²_4QD. Estimated from static (no jitter) 4-QD noise floor measurement. Typically R ≈ (0.5 μrad)².
Update step:
K_k = P_pred·Hᵀ·(H·P_pred·Hᵀ+R)⁻¹. x̂_k = x̂_pred + K_k·(z_k − H·x̂_pred). ³×³ matrix ops in FPGA fixed-point arithmetic.
Feed-forward cancel:
u_ff(k) = −K_ff·d̂_k injected before FSM DAC. Reduces RW harmonic disturbance by 15–25 dB vs. PID alone.
Notes: The velocity-form PID is critical to prevent integrator windup — if the FSM saturates at its travel limit, a standard PID would wind up the integrator and cause a step response on recovery. Velocity form automatically handles this.

## Slide 10
SOFTWARE — DSP & ANALYSIS
Signal Processing & Data Analysis Software
Digital Signal Processing Chain (Python / MATLAB)
Raw capture
12-bit digitizer streams I/Q or direct-detection waveform to PC via PCIe DMA. Typical burst: 10⁸ samples at 3 GSa/s = 33 ms record.
Clock recovery
Gardner TED (timing error detector) implemented in NumPy. Estimates sample clock offset and corrects by interpolation (poly-phase filter bank).
Carrier phase recovery
Viterbi-Viterbi 4th-power algorithm for QPSK; differential decode for DPSK. Compensates laser phase noise and Doppler shift.
Chromatic dispersion comp.
Frequency-domain equalization: H_CD(f) = exp(j·β₂·L·(2πf)²/2). Correct for fiber delay-line dispersion in fiber-emulated link.
BER counting / FEC
Error counting vs. known PRBS-31 pattern. FEC overhead: G.975.1 (7% overhead LDPC). Pre-FEC and post-FEC BER computed separately.
Optical Design & Simulation Software
Zemax OpticStudio
Geometric ray-tracing + physical optics. Sensitivity matrix generation. WFE budget construction. Export Zernike coefficients to Python.
COMSOL / ANSYS
FEA thermo-elastic analysis. Predicts mirror surface deformation under orbit ΔT. Feeds thermal WFE budget contributor.
MATLAB Control TB
Loop shaping, Bode plot, root locus. Notch filter design. Monte Carlo closed-loop stability analysis over plant uncertainty.
Python / SciPy stack
scipy.signal (filter design, Welch PSD), scipy.optimize (controller tuning), astropy (orbit geometry), pandas (telemetry logging).
Notes: The Gardner TED clock recovery is a standard telecom algorithm that works well for DPSK and OOK but needs extension for QPSK. The Viterbi-Viterbi phase estimator requires the laser linewidth × symbol period product to be < 10E-3 — this is why we specified < 100 kHz linewidth.

## Slide 11
INTEGRATED TEST CHAIN
Complete Hardware + Software Test Chain
Transmit side
AWG
(Keysight)
PRBS / IQ
data source
MZM
Modulator
OOK / DPSK
/ QPSK
EDFA
Amplifier
0.5–2 W
output
Beam
Expander
Galilean
40–80 mm
FSM
(PI S-330)
TX beam
steering
Free-space path / fiber delay-line emulator
ND attenuators + SMF-28 fiber spool (emulated range loss) + beam steering jitter injection
Receive side
Cassegrain
Telescope
50 mm
aperture
BPF +
Focuser
0.5 nm
passband
InGaAs
APD
M=12
−35 dBm sens.
TIA +
DSP
Demod +
FEC
BERT /
Oscilloscope
BER +
eye diagram
PAT tracking loop (closed)
4-QD
Detector
Hamamatsu
G6849
ADC
NI 9220
100 kHz
±10V
FPGA
cRIO
PID+Kalman
100 kHz loop
DAC
NI 9264
FSM drive
±10V
FSM
PI S-330
< 5 μrad
pointing err
Notes: This slide is the centerpiece — the complete integrated chain. Emphasize the three horizontal layers: TX chain (gold), RX chain (cyan), and PAT loop (teal), all operating simultaneously during an end-to-end link test. The fiber delay line emulates the free-space range.

## Slide 12
MEASUREMENT PROCEDURES
Key Measurement Sequences & Acceptance Criteria
WFE & Strehl Measurement
1
Mount optic on 6-DOF stage; inject collimated 1550 nm beam from TLS through system
2
Record Shack-Hartmann WFS data: 150-lenslet array, 100 frames averaged, Zernike decomposition via modal fitting
3
Compute RMS WFE from Z4–Z37 Zernike terms. Verify RSS total < 30 nm. Compute Strehl via Maréchal
4
Acceptance: S > 0.80. Flag contributors > allocation. Iterate polish / alignment until pass
✓ RSS WFE < 30 nm | Strehl S > 0.80
FSM Frequency Response & Loop Closure
1
Inject swept-sine ±0.5 mrad from AWG into FSM. Measure 4-QD output. Compute open-loop transfer function H(f) = Output/Input via FFT ratio
2
Fit second-order model P(s) to identify ω_n, ζ, K_act. Extract all mechanical resonance frequencies
3
Design PID + notch in MATLAB using loop shaping. Target crossover f_c ≈ 300 Hz, PM > 45°. Simulate closed-loop step response
4
Close loop on cRIO FPGA. Inject RW harmonic disturbance (60/120/180 Hz). Measure closed-loop error PSD. Verify RMS < 5 μrad
✓ Pointing error RMS < 5 μrad | BW > 500 Hz | Stable (no oscillation)
BER vs. Received Power (Waterfall Curve)
1
Setup: AWG → MZM → EDFA → VOA → APD → TIA → BERT. Fix data rate (10 Gbps PRBS-31)
2
Sweep VOA attenuation from 0 dB to −35 dB in 0.5 dB steps. At each step, record BER from BERT over 60-second integration
3
Plot BER vs. P_rx (waterfall curve). Fit to theoretical Q-function model: BER = 0.5·erfc(Q/√2)
4
Read sensitivity at BER = 10⁻¹². Compare to link budget requirement. Verify margin ≥ 3 dB
✓ Sensitivity ≤ −35 dBm @ 10 Gbps, BER = 10⁻¹²
Notes: These three procedures — WFE/Strehl, FSM loop closure, and BER waterfall — are the three most critical measurements that directly verify the design against the link budget and WFE budget requirements. Everything else is supporting characterization.

## Slide 13
ENVIRONMENTAL QUALIFICATION
Thermal Vacuum & Vibration Test Setup
Thermal Vacuum (TVAC) Test
Chamber:
1m³ stainless steel TVAC chamber. Base pressure < 10⁻⁶ mbar (turbomolecular pump). Temperature range: −100°C to +150°C (LN₂ shroud + heater elements).
Thermal cycling:
3–5 complete cycles per ECSS-Q-ST-70-04: cold soak at −40°C for 2h, hot soak at +70°C for 2h. Ramp rate: 2°C/min. Monitor WFE vs. temperature continuously.
In-situ WFE monitoring:
Fiber-pigtailed TLS + Shack-Hartmann WFS coupled through TVAC optical viewport (fused silica window). Records Zernike coefficients vs. temperature in real time.
Thermal control:
MLI blankets + thermoelectric coolers (TEC) on critical optics. Temperature telemetry: 16-ch NI-9213 TC module. Log ΔT gradient across aperture.
Pass criteria:
WFE RSS < 35 nm at all temperatures; Strehl S > 0.78 at thermal extremes; no optical element de-bonds or misaligns by > 10 μm.
Launch Vibration & Shock Test
Shaker system:
LDS V830 or IMV F-series electrodynamic shaker. Force: 18 kN. Frequency: 5–3000 Hz. 3-axis sequential testing (X, Y, Z).
Random vibration:
Spectrum per GEVS-7000A Table 2.4.1 (CubeSat level): 0.04 g²/Hz from 20–50 Hz, −6 dB/oct roll-off. Overall: 14.1 g_rms. Duration: 60 sec/axis.
Sine sweep:
5–100 Hz at 4 oct/min, 1.5g. Identifies resonant frequencies. Re-run post-random to check for frequency shift (> 5% = structural failure indicator).
Shock test:
Pyrotechnic shock: half-sine pulse 3000g, 0.5 ms. Simulates separation event. Post-shock: re-measure WFE and re-run BER test.
Instrument suite:
Accelerometers (PCB 353B15) on optical bench, telescope mount, and FSM. Data logger: LMS SCADAS. Post-processing: LMS Test.Lab or MATLAB.
Radiation:
EDFA Er-fiber and InGaAs APD irradiated to mission TID (10–50 krad Si) using Co-60 source. Monitor gain, NF, dark current vs. dose. Pre-anneal fiber at 100°C for 72h to establish stable degradation baseline.
Notes: TVAC and vibration testing are often the most schedule-critical part of a space hardware campaign. Plan for 6–8 weeks for a full qualification campaign. The in-situ WFE measurement through the viewport is the key differentiator — most programs only do before/after, but continuous monitoring catches transient misalignment events.

## Slide 14
SOFTWARE — DATA PIPELINE
Test Data Flow, Storage & Analysis Pipeline
01
Instrument acquisition
All instruments write to shared HDF5 file via NI-DAQmx streaming + custom Python SCPI logger. Time-stamped to GPS 1PPS reference. Typical dataset: 50–200 GB per full test day.
02
Real-time QA dashboard
Grafana + InfluxDB stack. Instruments publish telemetry via MQTT. Dashboard shows BER, WFE, pointing error, thermal map in real time. Alerts on out-of-spec events.
03
Post-processing pipeline
Python pipeline (pytest-based): load HDF5 → run signal processing modules → compute KPIs (BER, Strehl, pointing RMS, link margin) → compare vs. acceptance criteria → flag failures.
04
Statistical analysis
Monte Carlo over N=1000 test runs: compute yield (% of assemblies passing all criteria). Allan deviation of pointing error. Welch PSD of jitter. PDF fitting of WFE contributors.
05
Report generation
Automated Jupyter-to-PDF report: plots, tables, pass/fail matrix. Version-controlled in Git. Links each measured KPI back to design requirement. Traceable to link budget.
Key Python packages:
numpy / scipy
matplotlib
h5py (HDF5)
pyvisa (SCPI)
control (Bode/root locus)
astropy
pytest
jupyter
Notes: HDF5 is the ideal format for large test datasets because it supports hierarchical data, metadata, and partial reads — you don't need to load 200 GB into RAM to extract one channel. The Grafana real-time dashboard pays dividends during long thermal cycling tests where you need to catch anomalies immediately.

## Slide 15
SUMMARY
Test Setup Summary &
Readiness Checklist
Optical Bench
Isolation table (pneumatic active)
6-DOF kinematic mounts for all optics
Beam expander + polarization optics
Fizeau interferometer + Shack-Hartmann WFS
Calibrated ND filter set + power meter
TX Chain Instruments
Tunable laser source (Keysight 81600B)
Optical spectrum analyzer (Yokogawa AQ6370D)
Arbitrary waveform generator (Keysight M8195A)
VNA for EO bandwidth (Keysight E5063A)
EDFA gain + NF characterization setup
RX Chain & BER
InGaAs APD bias + characterization (Keithley SMU)
BERT (Anritsu MP1800A)
Variable optical attenuator sweep
Sampling oscilloscope (Keysight 86100D)
Electronic spectrum analyzer (Keysight N9010B)
PAT System
Jitter emulator (PI S-330 + AWG drive)
4-QD detector (Hamamatsu G6849)
NI cRIO-9049 FPGA controller
PI E-727 high-voltage FSM driver
CMOS InGaAs camera (acquisition phase)
DAQ & Software
PXI digitizer (NI PXIe-5186, 3 GSa/s)
GPS-disciplined 10 MHz reference
Python SCPI pipeline + HDF5 logging
Grafana/InfluxDB real-time dashboard
Automated pytest analysis pipeline
Environmental
TVAC chamber (< 10⁻⁶ mbar, ±100°C)
In-situ WFS through optical viewport
Electrodynamic shaker (GEVS-7000A)
Shock test (3000g half-sine)
Co-60 TID radiation facility
KAIST · Dept. of Aerospace Engineering · OISL Experimental Test Setup · 2026
Notes: The checklist maps directly to the six hardware/software categories covered in the deck. Use this as the procurement and lab readiness checklist. Each item should be confirmed calibrated and within its calibration interval before a test campaign begins.